Clearing old readings

At the beginning of each day, after the readings has been uploaded and posted, the neutron probe needs formatting to clear old readings.

| Press key | Read |  |
| --- | --- | --- |
| ENTER | READY | wake up the probe |
| FMT | REC 662 | space for 662 tubes |
| ENTER | KDATA=0 |  |
| ENTER | DPTHS=8 | confirm 0 keydata items |
| ENTER | SET FMT? | confirm 8 depths |
| ENTER | REC 662 | set format (space for 662 tubes) |
| ENTER | READY | back to ‘Ready’ mode |

The neutron probe is now ready for readings to be taken.

Access tubes are normally installed in groups of three, each tube having the same site number.  The site number should be written on the tube.  The readings from the three tubes will be averaged before posting.

Before posting the software will always pause to allow you to confirm the date and time.  You can also record the date (at K2) and time (at K1) as keydata items at a ‘dummy’ site with ID=0.

Taking readings in the field

Place the neutron probe on top of the aluminium access tube.

Carefully lower the detector to the last stop on the cable.

| Press key | Read |  |
| --- | --- | --- |
| ENTER | READY | wake up the probe |
| LOG | ID 662 |  |
| 123 | KDATA=0 | enter the site number |
| ENTER | TAKE 8 | and press Enter |
| START | COUNT 8 | starting a count... |
|  | M8 15123 | count = 15,123 |
| ENTER |  | press Enter lift the cable to next stop and then press Start |
| ENTER | DATA OK? | at top of tube... if data is ok press Enter |
| ENTER | READY | finished - go to next tube |

These steps assume that you are not entering any keydata items, and that the date and time that the readings were taken will be entered as the readings are posted.  503TDR probes (1998+) record the date and time automatically.

Research Services New England

T: (02) 9810 3563  F: (02) 9810 3323   <support@rsne.com.au>